## Slide 1
Automatic Multiple
Tritium Sampler
Notes: 기술개발_자동식 삼중수소 다중 시료 채취기-1

## Slide 2: CONTENTS
Ⅰ
Characteristics of CANDU Reactor & Tritium
Ⅱ
Tritium Measurements
Ⅲ
Needs of the equipment development
Ⅳ
Automatic Multiple Tritium Sampler(AMTS)
Ⅴ
Conclusions & Expectations
Ⅵ
Future Plans

## Slide 3
Characteristics of CANDU Reactor and Tritium
Ⅰ
  ▣ Natural Uranium Fuel contains 0.7% of U-235
  ▣ Use of the Heavy water (D2O) as Coolant and Moderator
▣ The interactions of neutrons with Heavy water
→ Tritium Occurrence
  ▣ CANDU Reactor in Korea → Wolsong Plant 1 & 2
Characteristics of CANDU Reactor

## Slide 4
▣ Hydrogen with 2 neutrons exists gas or vapor status.
  ▣ Pure Beta-emitting radionuclide, Half-Life : 12.3 years .
▣ Maximum Energy : 18.6keV , Average Energy : 5.6keV
  ▣ Internal exposure is the more harmful to the body via
  inhalation or wounds.
Characteristics of Tritium
Characteristics of CANDU Reactor and Tritium
Ⅰ

## Slide 5
▣ Use of the Fixed Tritium Air Monitors (FTAM)
Methods of Tritium Measurement
Tritium Measurements
Ⅱ

## Slide 6
▣ Use of the Portable Tritium Meter
Methods of Tritium Measurement
Tritium Measurements
Ⅱ

## Slide 7
▣ Use of the Liquid Scintillation Counter
Methods of Tritium Measurement
Tritium Measurements
Ⅱ

## Slide 8
▣ Use of the Fixed Tritium Air Monitors (FTAM)
FTAM is fixed and installed at the
position of specific area.
  Through 24 sampling ports lining with
  metal pipes, select one port
  to sampling by suction pump.
  Using the ion chamber detector
  to measure, Good reliance of
  Measurement .
Methods of Tritium Measurement
Tritium Measurements
Ⅱ

## Slide 9
▣ Use of the Portable Tritium Meter
  Easy to carry-on and comfortable
  to use.
  Simultaneously, suction and
  measurement of the sampling gas.
  Use of the ion chamber detector to
  measure, but reliance of measurement
  is normal.
Methods of Tritium Measurement
Tritium Measurements
Ⅱ

## Slide 10
▣ Use of the Liquid Scintillation Counter
With the similar chemical properties,
Tritium can dissolve into the water .
Air suction with the tritium → Bubbler
→ Sample bottle [100cc] →
Sample Volume for measurement [10L]
Methods of Tritium Measurement
Tritium Measurements
Ⅱ

## Slide 11
Collected tritium samples(1cc) and
fluorescent solution (10cc) prepare
by mixing
Measure the sample with
the scintillation counter
Using the scintillation detector to
measure, good reliance of
measurement.
Methods of Tritium Measurement
▣ Use of the Liquid Scintillation Counter
Tritium Measurements
Ⅱ

## Slide 12
Sampling | NPP#1 |  | NPP#2 |  | TRF [*2) | 
 | Positions | Times | Positions | Times | Positions | Times
Days | 10 | 30 | 9 | 27 | 6 | 18
Weekly | 3 | 9 | 3 | 9 | 2 | 6
▣ Observations of radiation for several areas in Radiation zone
with the routine [*1) tritium sampling and analysis
[*1) 3 shift operators on duty
[*2) TRF(Tritium Removal Facility)
Status of Tritium measurements
Tritium Measurements
Ⅱ

## Slide 13
  ▣ At the abnormal situation(Tritium concentration rising),
  add the sampling points around the suspicious leakage
  area.
  ▣ At the leakage situation of Tritium, add the additional
  sampling and measurements of Tritium concentration.
  ▣ One sample with a Bubbler.
Status of Tritium measurements
Tritium Measurements
Ⅱ

## Slide 14
▣ Lack of sampling equipment holdings for emergency
▣ No record sampling times, No time reservations features
▣ The unexpected change of flow meter can cause the
sampling volume.
Equipment
▣ Bubbler can sample only one point with a bubbler
  ▣ Sampling during abnormal concentration rising are
  difficult.
Methods
▣ A variety of sampling point & how to sampling by each
operator
Operator
Needs of the equipment development
Ⅲ

## Slide 15
Methods
▣ Bubbler can sample only one point with a bubbler.
▶ Sampling of multiple points need with a bubbler.
▶ During sampling at the same point, it needs
a equipment that able to automatic sample with
the setting time.
▣ During the abnormal situation , Operator increase
to sample because of the sampling interval time
is shorter.
Needs of the equipment development
Ⅲ

## Slide 16
▣ Busy Outage and equipment contamination can
cause a lack of equipment .
▶ Equipment needs multiple sampler line and
requires a self-cleaning function.
▶ Need the sampling time storage function .
▣ Equipment itself does not record the sampling time.
Equipment
Needs of the equipment development
Ⅲ

## Slide 17
Equipment
▣ The unexpected setting change of flow meter(LPM)
can cause the sampling volume change.
▶ Need the automatic correction features in
accordance with the change of flow rate(LPM).
▶ Need Sampling reservation function.
▣ No Sampling reservation function.
Needs of the equipment development
Ⅲ

## Slide 18
Main controller
Sample selector
Bubbler
Prototype of the Sampler
Automatic Multiple Tritium Sampler(AMTS)
Ⅳ

## Slide 19
Configuration of the Sampler
Automatic Multiple Tritium Sampler(AMTS)
Ⅳ
Solenoid valve
Operation Display LED
Control signals
Clean air
Bubbler
Sample selector
Main controller
Purge out

## Slide 20
Sample point 4
Sample point 3
Sample point 2
Sample point 1
Solenoid valve
Operation Display LED
Control signals
Work flow of the Sampler
Automatic Multiple Tritium Sampler(AMTS)
Ⅳ
Purge out
Clean air
Sample selector
Bubbler
Main controller

## Slide 21
8 Bit Microprocessor with Watchdog Timer
32 kbyte EEPROM, 8 kbyte SRAM
12 Bit A/D 4 Channel, 12 Bit D/A 2 Channel
8 key Input (cursor movement, decrease, return, action)
240 x 64 Pixel Graphic LCD with LED Back Light
Real Time Clock Chip (Back Up Battery life 10 years)
10 Port Samples Selector Control (Power transistor driven by the relay)
Bubbler Control of Photo Coupler
(work orders, stop orders, work situation )
Main Controller
Specifications
Automatic Multiple Tritium Sampler(AMTS)
Ⅳ

## Slide 22
4 Input Port + 1 Clean Air Port, 4 Out Port + 1 Purge Air Port
Path selection by solenoid valves (DC + 24V powered)
Power transistors, driven by a relay (10 Port)
Sampling Air Selector
Photo Coupler automatically controlled by the Bubbler
  Subject made by the purge, stabilization, bubbling, automatically
  perform a clean operation
10 Liter Air Sampling (3 Liter Per Minute Sampling Speed)
Bubbler
Specifications
Automatic Multiple Tritium Sampler(AMTS)
Ⅳ

## Slide 23
AMTS Test 1 – Continuous Measurements
  ▣ The time required for 4 points sampling with one Bubbler
(Sampling time = 10minutes / sampling x 4sampling = 40minutes)
  ◈ The necessary time applied with AMTS
(Sampling time = 6minutes / sampling x 4sampling = 24minutes)
▶ 16 minutes shorter
Application Tests at Local Point
Automatic Multiple Tritium Sampler(AMTS)
Ⅳ

## Slide 24
AMTS Test 1 – Continuous Measurements
  ▣ The time required for 4 points sampling with two Bubbler
  (Sampling time = 10minutes / 2bubbler x 4sampling = 20minutes)
  ◈ The necessary time applied with AMTS
  (Sampling time = 6minutes / 1bubbler x 4sampling = 24minutes)
▶ Two bubbler effects
Application Tests at Local Point
Automatic Multiple Tritium Sampler(AMTS)
Ⅳ

## Slide 25
AMTS Test 2 – Time settings
▣ To reduce Sampling Error – Sampling time
(Time on/off by operator – human error)
◈ AMTS sampling program
- Sampling start at scheduled time ( 30 minutes, 1 hour …)
→ Automatically purge and self-cleaning
▶ Realization and reliability of unmanned automated collection
Application Tests at Local Point
Automatic Multiple Tritium Sampler(AMTS)
Ⅳ

## Slide 26
Save the Sampling time
With the developed equipment,
it shows
two equipment effects
with one bubbler.
Standardization
Of Sampling Procedures
Anyone who has a
certain time can sample
in the same way.
Unmanned automation
After the initial setup,
all sampling can obtain
automatically without
people.
Also, operators can do any
other tasks and services.
Unmanned automation and standardization of Sampling depending on reliability improvement
Unmanned
automation
Save the
Sampling
time
Automatic Multiple Tritium Sampler(AMTS)
Ⅳ
Standardization
Of Sampling
Procedures

## Slide 27
When tritium sampling at the several area in Radiation Control zone,
it expects the shorter time to sample with the AMTS.
More then 2 sampling positions can concentrate a sampling position.
The more effects the tritium concentration surveillance .
With the reservation function of AMTS, it is possible to sample and to purge periodically according to the setting time and the setting procedure.
Conclusions & Expectations
Ⅴ

## Slide 28
When heavy water leakage breaks out, it can sample on the various areas within the shorter time, takes any actions quickly for emergency situation.
After the outage, it is possible to observe quickly around the suspected heavy water leak area during the reactor start operation.
After the outage of Wolsong Unit #1, it is possible to observe quickly around the suspected heavy water leak point.
Conclusions & Expectations
Ⅴ

## Slide 29
AMTS has a sampling pump ,
sampling can not be done more than 2 points simultaneously.
Consists of the three separated elements,
it is difficult to carry on and movement.
At the abnormal loss of power,
it can be lost all of sampling data.
Improvements
Simultaneous
function
Portable
function
Storage
function
Future Plans
Ⅵ

## Slide 30
Simultaneous function
Improve the equipment to sample more than 2 points at the same time.
Portable function
Improve the equipment All- in-one and portable.
Future Plans
Ⅵ

## Slide 31
TAM function
Apply TAM function to portable tritium measurement.
⇒ Unmanned automatic tritium measurement.
Storage function
Improve the program to store the sampling status automatically.
Future Plans
Ⅵ

## Slide 32
Thank you